## Cover
Department of Veterans Affairs Requirements Traceability Matrix
Business Requirements Matrix Anesthesiology EHRM Perioperative Council
January 2024

## Executive Summary
EHRM Business Needs Summary
This document provides an overview of the business needs submitted on behalf of the VA National Anesthesia Program to the EHRM Perioperative Council..
Title
EHRM Anesthesia (WC20200807)
Points of Contact
Work Group | Anesthesiology and Critical Care
Work Group Lead: | Dr. Art Wallace Dr. Roberta Reedy
Key VHA SMEs: | Art.Wallace@va.gov San Francisco Chief of Anesthesia Roberta.Reedy@va.gov Seattle CRNA Christopher.Kelley3@va.gov Spokane CRNA <Enter name and email address(es) of key VHA SMEs who can speak to the requirements, if needed> NOTE: Cerner will use the SMEs identified in this field to form an inclusive list of SMEs they can discuss enhancement with when developing
Requirements Support Team: | Rhonna Clark <Rhonna.Clark@va.gov> Emily Jevec <Emily.Jevec@va.gov>
Cerner POCs: | NA
Overview
Current Processes
VHA Anesthesiology Professionals must provide safe and efficient anesthesia and critical care services for over 500,000 procedures each year, and must include Quality Improvement initiatives and evaluate clinical performance measures by using electronic anesthesia record keepers (ARK) and critical care information systems (CIS). VHA currently has ARK/CIS systems fully operational at all VHA medical facilities for over a decade and are a proven reliable, safe, and efficient system for providing care that are designed and optimized specifically for the complex and high acuity of veteran-specific care.
Description of Issue
The Cerner SA Anesthesia solution was deemed by the National Anesthesia Program and Anesthesiology Subject Matter Experts (SMES) as unsafe and unfit for clinical use.Concerns were based on comprehensive five-year review by VHA Anesthesiology SMEs. A comprehensive defect report was created and none of the issues have been remediated over these five years. Despite five years of continuous feedback, there has been no remediation of the VA requested modifications to the SA Anesthesia solution, which include: increased risk of error in medication administration and charting, decreased visibility in physiological monitoring and situational awareness, and significant system usability issues as well as increased dependency on manual, paper-based processes and information charting. Deployment proceeded despite subject matter expert advisement that moving to the Cerner SA Anesthesia solution presents a high potential for dangerous errors, is not safe for clinical use, and is unable to provide safe patient care at this point for go-live. Deployed Cerner SA Anesthesia does not meet the outlined requirements for VA Anesthesia to deliver safe anesthesia care at the level offered by the current VA Anesthesia Record Keeper Systems. Cerner SA Anesthesia does not meet most of VA's IT Anesthesia needs, VA minimum specifications to meet current VA functionality, and does not resolve usability, configuration, and customization requested during council discussions. The Cerner SA Anesthesia solution has to be optimized, fast, easy to use, easy to read, and provide complete information. The product must meet at a minimum the current Anesthesia Record Keeper standards (addressed under Nat'l Mandatory Elements, Required ARK Elements, Required CIS Elements, and Required Surgery Pkg Elements) to enable the ability to evaluate clinical performance measures and provide the cadre of detailed reporting and analytics to maintain Anesthesia Program's abilities to improve outcomes for patients; as well as the requirements on the Requirement Traceability Matrix (RTM) which identify the critical patient safety issues that need to be addressed by Cerner solutions, and resolve serious usability concerns and needs. Through testing, SA Anesthesia demonstrated issues that present Patient Safety risk: • Inefficient and inaccurate medication charting requires more action to accomplish tasks than current VA capabilities and impacts focus on Patient care • Incomplete drug library with poor search functionality to quickly document medication administration • Inaccurate or inappropriate dosing and concentrations presenting Patient safety risks • Poor usability and readability affecting ability to ensure correct anesthesia administration, documentation, or monitoring and present high risk of loss of situational awareness • Poor display and readability is incompatible with maximum Personal Protective Equipment (PPE) required for pandemic protocols (e.g. COVID-19) in the operating room • Inefficient capture of the transfusion process in anesthesia record • Lack of automation requiring Anesthesia Providers to conduct a number of procedures on paper and with manual data entry presenting Patient safety issues • Lack of support in the creation of many standardized processes via Power Plans and Anesthesia-specific notes that meet VA specifications
Value of Change
The Cerner Surginet Anesthesia (SA) has been deemed unsafe for clinical care in the VHA by the designated EHRM Anesthesaiology staff and SMEs. The Cerner SA solution presents high risk for life-threatening errors in the perioperative environment, from preop, intraop, and postop care. GIven lack of remediation of hundreds of defects identified, and lack of compliance with the core National Mandatory Elements, a safe, non-disruptive approach to resolving this was through interfacing Cerner to the existing operational CIS/ARK systems, which meet all of the National Mandatory Elements today, and have been proven through operation for over a decade at all VAMCs that provide anesthesia and critical care.
Escalation Efforts To Date
VHA Anesthesiology and Critical Care National Mandatory Requirments was provided to EHRM Perioperative Council (see National Mandatory Elements). The Anesthesia Working Group members and discussed with Cerner throughout Phase 1 Workshops and EHRM Perioperative Council meetings
Associated Business Requirements and Architecture (with Repository IDs) Note: the Business Process Models (BPM) identified below are Visio files. Information on downloading Visio Viewer (OIT approved) to view the BPMs is located on the last tab.
Requirements Traceability Matrix |  | Located in separate spreadsheet
Work Collection (Supporting Information) |  | WC20200807_EHRM Anesthesiology <https://dvagov.sharepoint.com/sites/VHARDM/Deliverables/WC20200807%20Anesthesia>
VHA Business Process Model requirements |  | BPM3285 -- Provide Anesthesia Record Keeping (ARK) Monitoring, As-Is BPM3286 -- Provide Anesthesia Services, As-Is BPM3287 -- Provide Pre-Anesthesia Care, As-Is BPM3288 -- Provide Anesthesia Monitoring, As-Is BPM3289 -- Provide Post-Anesthesia Care, As-Is

## National Mandatory Elements
Ref. | Category | Sub-category | Mandatory Element | Mandatory ICU | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op) | Offeror Planned (P)/Met (M)/Unmet (UM) | Descriptions | Comments
1.0 | National System Requirements
1.1 |  | National Database Implementation | Must Demonstrate following Requirements Prior to Contractual Obligation with VA or VISN | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.1.1 |  |  | The database must be identical within all hospitals in the VA using this ARK or CIS system. All database and template modifications must be entered only once and immediately propagated to all VA hospitals using that ARK or CIS system. Individual hospital databases and templates must be locked to prevent local modifications. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
|  |  | The parameters recorded in the ARK or CIS system will be by national mandate. Local facilites or vendors or users may not turn off nationally mandated parameters either for all systems or for individual patients. Parameters recorded must be locked and may not be changed by the user. A national standard terminology will be used for recorded parameters, medications, and events. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.1.2 |  |  | Vendors must provide a mechanism whereby nationally-directed changes in the database and/or user templates can be pushed out to all facilities throughout the VA. A mechanism(s) must be provided to prevent changes to flow sheet template and database elements that might be mandated by the VA nationally for purposes of standardization. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.1.3 |  |  | The ability to create new database elements, controlled at the national level, must be maintained. Note: this must not prevent the ability to assign informational lines (e.g., specific templets or assessments), unique for the management of a specific patient, for use at the individual facility level and to be included in the patient’s permanent medical record. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.2 |  | Security and Reliability | Systems must meet all VA security requirements as well as ensure data against system failures (disaster recovery) and user input or errors (audit trail). | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.2.1 |  |  | Reliable and robust application able to survive network and server failures without loss of data or of real time data display. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.2.2 |  |  | Provide plans to ensure CIS and ARK function, production and database server redundancies for hospital/VISN/IT region including network failures. In the event of a network failure the CIS and ARK system should continue to function, record data, and allow users to log in, select patients, and use the system on a local basis until the network function recovers. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.3 |  | Data Extracts | Must Demonstrate following Requirements Prior to Contractual Obligation with VA or VISN | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.3.1 |  |  | A data extraction system for ALL clinical and administrative data in the ARK or CIS database must be capable of creating data extracts at least every 4 hours without degrading performance of ARK or CIS. Medications, physiologic parameters, times, events, notes, must be extractable in usable and exchangable data format. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.3.2 |  |  | The files must be in XML, ASCII flat files, or HL7 format. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.3.3 |  |  | Real time data from patient monitors, infusion pumps, ventilators, assist devices, etc. must be stored and extracted from the ARK or CIS at frequencies independent of the display time period and, in no case, less than every 1 minute.
1.4 |  | Data Standardization |  | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.4.1 |  |  | ARK and CIS data must meet all VA nomenclature and standardization requirements. The ARK and CIS database extracts must be mapped to a single standard VA-wide concept and mapped to the VA Unique Identifier (VUID) in VistA such that all HL7 interfaces map correctly to the same concepts in VistA. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.5 |  | VistA Interface Functionality | Must Demonstrate following Requirements Prior to Contractual Obligation with VA or VISN | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.5.1 |  |  | All required interfaces (see Interface Section, 4.0) with VistA must be certified prior to contract signing and installation in VA alpha site | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.6 |  | Remote Access and Off-line Viewing | ARK and CIS application must support remote view capabilities within and from outside hospital utilizing VPN and Citrix | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.6.1 |  |  | One or more clinically active ARK or CIS records must be able to be viewed entirely by authorized users or administrators on remote, non-clinical workstations throughout the local area network (LAN). Access of ARK and CIS records by remote, non-clinical workstations must be throught the thin client as well as via a browser-based application or a special viewing program. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.6.2 |  |  | All off-site performance monitoring of ARK or CIS hardware or software must satisfy VA Cyber-Security Requirements, the Administrative Simplification provisions of the Health Insurance Portability and Accountability Act of 1996 (HIPAA), and the provisions of The Privacy Act of 1974 and its Veterans Affairs applications. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.6.3 |  |  | Upgrades to the ARK or CIS server or workstation software that are accomplished by the manufacturer through remote access must satisfy VA Cyber-Security Requirements, preserve firewalls, and meet Enterprise Architecture Standards and VA security guidelines for virus protection. Software or hardware upgrades must not disrupt or alter a clinically active ARK or CIS record. An audit trail of all software changes must be maintained. All upgrades must be coordinated with COR. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.6.4 |  |  | Portable clinical ARK and CIS workstations must be fully functional in remote, non-operating room applications using local area network (LAN) connections to the ARK or CIS server. It must remain capable of importing physiologic data from its associated monitoring devices and anesthesia delivery machine. |  |  | Mandatory (ARK)
1.6.5 |  |  | CIS or ARK records that are not clinically active and have not yet been archived or exported to VistA must be able to be annotated by authorized users and administrators specifically granted privileged access to do so. Such annotations must be able to be performed at non-clinical, remote workstations that satisfy VA Cyber-Security Requirements, the Administrative Simplification provisions of the Health Insurance Portability and Accountability Act of 1996 (HIPAA), and the provisions of The Privacy Act of 1974 and its Veterans Affairs applications. An audit trail of changes, additions, and annotations to the record must be maintained. |  |  | Mandatory (ARK)
1.8 |  | Windows desktop compatibility | ARK and CIS application must not conflict or interfere with the function of any VA software installed on any Windows clinical workstation either locally, or remotely (Citrix remote desktop). If so the conflict must be identified and remediated. These Windows desktop applications include, but are not limited to, the following: • BCMA • VISTA/CPRS/Protoviewer • VISTA Imaging • Clinical Note Templates • Internet Browsers (Firefox, Chrome, Edge) • Microsoft Offices Products • Java/Active X-Plug Ins • SMS • McAfee Virus Scan • ISS Proventia firewall • Sentillion Vergence Locator • Analytic Database • The CIS client application Microsoft Terminal Services or Citrix should be supported for remote viewing within the VA Firewall and the via One VA VPN. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.9 |  | Implementation |  | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.10 |  | Maintenance program | A uniform, multi-year maintenance program must be available for the VA, facilities, and or VISNs and must include all application upgrades, new versions, patches, etc. This program must ensure that all VAs are using the same ARK and CIS version. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
1.11 |  | Device Drivers | Device drivers to all VA medical devices for patient use in OR, ICU, PACU, etc. must be capable of accepting all data, measured and calculated. These drivers must be available without additional cost to all VA facilities. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
1.11.1 |  |  | The CIS/ARK must be capable of importing physiologic data as a continuous stream from most standard patient monitoring devices. This data must include, but will not be limited to, invasive and non-invasive blood pressure, heart rate, oxygen saturation, end-tidal carbon dioxide concentration, fraction of inspired oxygen, respiratory rate, and temperature. When available, inspired nitrous oxide concentration, fraction of inspired air, cardiac output, pulmonary artery pressures, central venous pressure, and ST segment analysis must be included. The CIS/ARK must be capable of accepting input data from standard United States-specification anesthesia delivery machines. This data will vary, depending upon the age and design of the anesthesia machine, but must include ventilator settings, gas flow rates, and anesthesia concentration settings when available. The recorded parameters may not be deselected or turned off by the user. Standard physiologic parameters, alarms, ventilator setting, anesthetic machine parameter, error codes will be recorded at least every 60 seconds and may not be turned off or deselected by the user. |  |  | Mandatory (ARK)
1.12 |  | Device Features | Once properly accessed, ARK and CIS workstations must have the capability to toggle (foreground/background continuous functionality) between the ARK and CIS applications and the Computerized Patient Record System (CPRS). | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
2.0 | System Requirements and characteristics
2.1 |  | Interoperability | Must Demonstrate following Requirements Prior to Contractual Obligation with VA, VISN, or facility.
2.1.1 |  |  | ARK and CIS applications in every clinical monitoring location [ICU, PACU, GI, Cardiology, NORA locations, and OR] must provide interoperability capabilities by either one of two methods below: A) By joining data from diverse sources through capabilities to interface with an Analytic Database, sending images to VistA Image and using a 'remote view' mode from the alternate CIS, or ARK or B) By an automatic and seamless exchange of data within a hospital 1) to permit clinical review of all data at any location in the hospital, 2) to provide integration of continuing care data (fluids, drips, medications, etc.), and 3) to allow complete analyses of clinical care. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) |  | Needs to be made strictor. This allows PDF to solve all problems.
2.2 |  | Workstation Requirements for CIS and ARK | Applications must be able to be run on the typical VA windows workstation hardware as listed below. Any requirements that deviate from this standard must be clearly outlined. Contractors will however be required to submit a list of equipment required which will include the minimum requirement for the ARK or CIS and Interface Software to operate at an efficient level as intended. The list of this equipment (including estimated pricing for each individual item suggested) shall be provided with this signed proposal and shall be included as defined in D219. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op) | Do we want to say Windows, everything else in the VA is Windows.
2.2.1 | IT Requirements and characteristics |  | Although each operating room has individual temperature and humidity controls, CIS/ARK workstations must be fully operational in an environment that is between 4 and 40 degrees Celsius (40 – 104 degrees Fahrenheit) and up to 100% humidity. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
2.2.2 |  |  | Operating rooms often contain surgical devices that are capable of generating unacceptable electronic artifact in concurrently used susceptible equipment. All CIS/ARK workstations must be acceptably free of electronic artifact when common surgical devices, such as electrocautery or surgical lighting, are in concurrent use. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
3.0 | Network Requirements
3.1 |  |  | ARK and CIS client application workstation shall be able to be run with a standard 100 full duplex network connection to the ARK or CIS server. The Contractor is required to submit with his/her signed proposal, a list of any additional network resources that are required for any associated ARK or CIS connection. The above must be clearly outlined and included with the Contractor’s response as Attachment D, Exhibit 6. The VA shall provide all necessary network hardware to meet this connection speed. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
3.2 |  |  | The ARK or CIS Contractor shall provide a listing of port and protocols used by the Ark or CIS system and include in Attachment D, Exhibit 7. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
3.3 |  |  | The ARK or CIS software shall be able to be run on PC located on the hospital domain. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
3.4 |  |  | The ARK or CIS system shall be able to operate over a WAN with a T3 connection. Contractor shall ensure that VISN or National Anesthesia Service (NAP) will be able to operate servers at remote locations for decreased hardware maintenance. A diagram of the VISN and National Anesthesia Service (NAP) WAN is included in this RFP. (See Tab - WAN Diagram) | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
3.5 |  |  | The CIS software shall have ability to operate off a WAN link with common intra VISN network latencies of up to 30 mS and VPN latencies of up to 100mS | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
3.6 |  |  | Reliable and robust application able to survive network and server failures | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
3.7 |  |  | Provide plans for ARK or CIS production and function and database server redundancies for hospital/VISN/IT region including network failures. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0 | VIistA Interfacing Requirements
4.0.1 |  | a. Medical Device Interfaces (HL7 and General Medical Device Interfaces | VHA will provide a uniform interface across all VISNs and hospitals to share information between VistA and ARK/CIS used for the ICU, ARK, PACU, GI, NORA, and Anesthesia Preop. The interface primarily will use HL7 communications as well as VistA Broker Calls (APIs). The vendor side of the interfaces must be developed, tested, implemented and supported according to VHA standards provided by Document Storage Systems (DSS), 12575 US Hwy 1, Suite 200, Juno Beach, FL 33408, 561-227-0207. DSS will provide programming specifications and development support upon request.
4.0.1.1 |  |  | 1) ARK or CIS device drivers or other electronic communication means to all VA medical devices for patient use must be capable of accepting all data, measured and calculated. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.2 |  |  | 2) Provide HL7 (or other) interfaces to all VISN XX site physiologic monitoring systems for VISN XX including: | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.2.1 |  |  | a) GE | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.2.2 |  |  | b) Philips | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.2.3 |  |  | c) Draeger
4.0.1.2.4 |  |  | d) Picis | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.3 |  |  | 3) Provide connectivity the following VISN XX medical devices for VISN XX, such as (see also Attachment F): | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.3.1 |  |  | a) Infusion Pumps | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.3.2 |  |  | b) Critical Care Ventilators | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.3.3 |  |  | c) Anesthesia Machines | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.3.4 |  |  | d) IABP devices | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
e) Medical Monitors, cardiac output monitors, pulse oximeters, cerebral oximeters, EEG monitors
f) Cardiopulmonary bypass machines, ECMO machines,
4.0.1.3.5 |  |  | g) Dialysis/CRRT Machines | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.4 |  |  | 4) Additional medical devices interfaces may be required as new equipment is identified for interfacing existing equipment/software identified herein for the CIS. These medical device interfaces will be developed and provided at no charge to the VA upon request of the VISN XX COR. If however, interfaces are required for those excluded from the above and additional charges may apply, the Contractor shall provide a written proposal to include a detailed description of the product and estimated pricing structure associated with any additional interface with this proposal as Exhibit 9 The above will be reviewed during the evaluation process however costs for the above are not to be included in the final price for this contract. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.1.5 |  |  | 5) Applications collection data from these medical devices shall run as a service or a daemon and shall not require a user to be interactively logged into the computer to collect the data. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.2 |  | b. VistA Interfaces Government Furnished Property
4.0.2.1 |  |  | Overview: The Contractor must provide interfaces to solve the problem of electronic documentation, continuum of care, advance clinical access, and clinical workflow in critical care areas. The interfaces will exchange data bi-directionally between VistA/CPRS and ARK/CIS applications. The interface is a key element in the creation of an electronic medical record for the critical care areas. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.2.2 |  |  | The Government will furnish the appropriate interfaces to operate on the VistA systems. These interfaces are referred to as the VistA-side interfaces. That is, they provide only one half of the interface requirements. The successful Contractor will be responsible to provide the other half, (the vendor-side) as part of this contract. The Contractor shall coordinate the above with the Primary COTR and will work with either VA staff and/or third party Contractor to make a successful connection and transfer of data with the VistA system through the Government provided interface. If costs are incurred by the third party to properly make the necessary interface connection(s) operate correctly with the products provided by this contract, it is the Contractor’s responsibility to cover those costs. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.2.3 |  |  | For this particular contract and the products being acquired, the VistA-side interfaces have been developed by a third party. That company is Document Storage Systems (DSS), 12575 US Hwy 1, Suite 200, Juno Beach, FL 33408, 561-227-0207. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.0.3 |  | c. Description of Integration Requirements
4.0.3.1 |  |  | The functional integration requirements needed to meet VHA requirements are listed below. The VA-side of the VistA interface will be provided to the successful Clinical Information System (CIS) Contractor for the ICU CIS and or OR ARK. The ARK/CIS Contractor will be required to have a VistA testing environment, QA staff, and provide installation, training and support of the ARK/CIS interface. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1 |  | d. Vendor Side Integration
4.1.a |  | User Authentication (Security) — For every sign on vendor software will use the VA master list to validate users and user rights for the clinical application through VistA Broker Calls.
4.1.1 |  |  | 1. Purpose: To provide a single sign-on source for permitting all users access to CIS and/or ARK using only VA user access/verify codes. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1.2.1 |  |  | a. Users will sign-on to CIS and/or ARK using their VistA access/verify codes. The logon dialog box must mask (e.g. asterisk ***) both access and verify codes. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1.2.2 |  |  | b. A separate authentication system is not maintained in the CIS and/or ARK for standard logons. A shadow list (access/verify/signature codes) may be employed for use only when the network or VistA is down utilizing VA security standards. All changes in VA access/verify codes must be used immediately for subsequent CIS and/or ARK authentications. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1.2.3 |  |  | c. A successful logon through VistA verifies the user name and VistA credentials. This Vista information must match the information displayed in the CIS and/or ARK. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1.2.4 |  |  | d. The users' menus/keys, allocated in VistA will determine users’ privileges within the CIS and/or ARK. The CIS and/or ARK application will allow creation of a “user group” and assign user groups specific levels of CIS and/or ARK access, such as 'read/write', 'read-only', and specific access privileges. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1.2.5 |  |  | e. The Vista interface determines the user group privileges using “rules” based upon the user’s keys assigned in VistA. Keys will be created in the CIS and/or ARK to match the Vista user group. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1.2.6 |  |  | f. Rules defining user groups are included with this interface and individual VISNs may request that additional Keys be created to meet special documentation needs; i.e., students and residents. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1.2.7 |  |  | g. If the logon is repeatedly unsuccessful, the user may not gain further access to the application. Logons will be limited to a fixed number of failed attempts defined by the facility. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1.2.8 |  |  | h. A temporary method must only be available to allow users to logon if network access is not available. As soon as the network is available, and it is clinically feasible—user must logon to the CIS and/or ARK by the standard method. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.1.2.9 |  |  | i. Electronic Signature data will be sent to CIS and/or ARK to allow user signing in CIS and/or ARK, including notes and records signed in prior to sending to VistA. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2 |  | Patient Selection / Registration in CIS — Individual patients will be selected by methods conforming to VistA standards, minimizing data entry, and by a process similar to patient selection in CPRS. Note: the selection of patients must occur through a VistA query utilizing VistA Broker calls (APIs).
4.2.1 |  |  | 1. The Master Veteran Index (MVI) D88must be used to accurately select individual patients and send patient specific information to the CIS and/or ARK with updates triggered by the CIS and/or ARK and VistA as necessary. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.1 |  |  | A patient search will be performed using a supplied API. The information returned will mimic the information displayed in CPRS. Or spell out wards units etc. on census lists and bed locations. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.2 |  |  | Patient name and VA unique patient identifier number must be prominently displayed on every page of the electronic display and record images. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.3 |  |  | Access to a Sensitive Patient Record will mimic CPRS and/or the current GUI. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.4 |  |  | a. Patients will be selected after the search is preformed from within CIS and/or ARK by multiple methods including the following: | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.5 |  |  | i.The contractor shall display at least but not limited to two unique patient identifiers including full name, social security number, date of birth and/or gender. Patient identification information consists of multiple identifiers and the CIS and/or ARK patient identification process must meet VA standards. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.6 |  |  | ii.a ‘John Doe’ option permitting use of the CIS and/or ARK without first requiring patient registration in VistA. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.7 |  |  | There must be a reliable method to retroactively match duplicate identities stored within CIS and/or ARK, and VistA systems with ‘John Doe’ or temporary patient identity to prevent the creation of duplicate patient records in CIS and/or ARK. Need IT, safety and Sara Temlitz to review. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
|  |  | There must be a reliable method to remove erroneous records from patient files. If an incorrect patient identifier is used and data is collected on the wrong patient there must be a way to correct the record and transfer the record to the correct patient. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.8 |  |  | There must be a reliable method to retroactively merge records stored within CIS and/or ARK, and VistA systems to create a continuous patient record. This must be permission based only. Need IT, safety, Stephania Griffin, and Sara Temlitz to review. Validate across all vendors. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.9 |  |  | c. Upon registration of a valid patient in the CIS and/or ARK, the patient information will be sent to CIS and/or ARK triggered by multiple events including query from CIS and/or ARK, ADT message, surgery schedule message or patient information update. The patient information, to be updated in CIS and/or ARK, will include, not limited to, patient identification identifiers, patient demographics, Code Status, CWAD(Crisis Warnings and Advanced Directives), bed location, treating specialty, as defined by transaction type. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.10 |  |  | i. New messages to CIS and/or ARKare sent for updates to the patient data stored when MVI updates data in VistA. CIS and/or ARK+D177 must update its database with the new data contemporaneously. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
4.2.1.11 |  |  | ii. CWAD will accompany patient registration/ADT messages, as well as provide a ‘re-send message’ trigger when information is updated. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.2.1.12 |  |  | iii. If available for each patient, CIS and/or ARK must accept VISN/hospital specific team names, ICU locations, OR and NORA locations (e.g., bed names) and treating specialty names, (e.g., MICU, SICU, CCU) to correctly identify patient locations and clinical provider groups/teams. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
4.2.1.13 |  |  | d. Contractor must provide the ability to allow the emergent, manual entry of patient demographics into the CIS prior to Registration through ADT in VistA. Identification and resolution of any differences once VistA registration is accomplished must be facilitated by CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
4.3 |  | Patient Status Information — Allergies and Code Status information is sent with initial registration of patient in CIS and with all updates to Status content.
4.3.1 |  |  | A) Allergies — Current allergies in VistA are in separate messages. A complete allergy set will be sent with a CIS query, ADT messages or surgery scheduling (SR) accompanied by pharmacy orders (ORM) based upon actions within the CPRS, ADT or surgery scheduling system, as well as patient registration in CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.3.2 |  |  | 2. Purpose: To provide allergy data stored in VistA for display in ARK or CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.3.3.1 |  |  | a. A read-only set of allergy data is sent in AKR or CIS for display-only. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.3.3.2 |  |  | b. Allergy messages are triggered by event changes managed in VistA. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.3.3.3 |  |  | c. Allergy updates made in VISTA must be updated automatically and in real-time in the CIS. Changes must be displayed in CIS both as a 'new' or 'critical' data alert and in an updating data field. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4 |  | Patient Information —In general, patient data are sent based upon 1) on a process external to the package (‘trigger’), 2) an update or addition to the package information, or 3) a query generated by the CIS. Queries may have multiple parameters set to define ranges of interest and typically utilize VistA Broker Calls (APIs). (NOTE: Where approriate, all patient data must be displayed on timelines in the CIS at the correct times in graphic and tabular formats.)
|  |  | D) Lab Orders — All lab orders will be sent to ARK or CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
|  |  | 1. Purpose: Lab orders will be sent to facilitate nursing schedules in CIS or ARK. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
|  |  | a. Lab orders will be sent to CIS for display. Including but not limited to blood bank and microbiology. This information should match CPRS or the current GUI. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
|  |  | b. All current orders including modifications will be made in VistA and contemporaneously displayed in CIS. Changes must be displayed in CIS both as a 'new' or 'critical' data alert and in an updating field. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.1 |  |  | A) Lab Results — All lab messages will be received and maintained in the CIS and/or ARK including microbiology. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.1.1 |  |  | 1. Purpose: To send including but not limited to: hematology, microbiology, blood bank, all satellite laboratory, pathology, and chemistry lab results to the CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.1.3 |  |  | a1. Hematology and chemistry lab results will be sent from VistA Laboratory package to CIS for display. No limitations can be placed on maximum number of laboratory data (specific tests or numbers of test results). | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.1.4 |  |  | a2. Microbiology, pathology, and blood bank, including entire culture reports and antibiotic sensitivities, and must be displayed. All parent / child reports must be included with each update. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.1.5 |  |  | b. ‘Out of range’ and ‘critical value’ labels will accompany results as appropriate. CIS must accept and display these data. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.1.6 |  |  | c. Results formats will be modified to permit CIS to display logical categories (e.g., CBC panel) rather than just alphabetical order. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.1.7 |  |  | d. Query capabilities initiated by CIS will mimic CPRS (date range, specific results, etc.) | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.1.8 |  |  | e. Changes must be displayed in CIS both as a 'new' or 'critical' data alert and in an updating data field. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.2 |  |  | B) Pharmacy Medication Schedules and Administration — All medications will be sent to ARK and CIS including include both outpatient and inpatient medications. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.2.1 |  |  | 1. Purpose: To send medications schedules and administration to the CIS for display-only. | Mandatory(ICU)
4.4.2.4 |  |  | a2. Inpatient administration events contained in BCMA will be sent contemporaneously to CIS for read only use | Mandatory(ICU) |  | Desirable
4.4.2.5 |  |  | b. Verified and cancelled order messages will be included. | Mandatory(ICU)
4.4.2.6 |  |  | c. Queries capabilities initiated by CIS will mimic CPRS (date range, specific results, etc.) | Mandatory(ICU)
4.4.2.7 |  |  | d. All order modifications will be made contemporaneously in VistA and CPRS and/or current GUI D162 | Mandatory(ICU)
4.4.2.9 |  |  | f. Changes must be displayed in CIS both as a 'new' or 'critical' data alert and in an updating data field. | Mandatory(ICU)
4.4.3 |  |  | C) Nursing Orders — All Nursing Orders will be sent to CIS in delimited form when available (e.g., restraint orders, pressure ulcer, fall risk). | Mandatory(ICU)
4.4.3.1 |  |  | 1. Purpose: Nursing orders will be sent to provide treatment schedules in CIS. | Mandatory(ICU)
4.4.3.3 |  |  | a. Nursing orders will be sent to CIS for read-only use. | Mandatory(ICU)
4.4.3.4 |  |  | b. All nursing order modifications will be made in VistA and contemporaneously displayed in CIS. Changes must be displayed in CIS both as a 'new' or 'critical' data alert and in an updating field. | Mandatory(ICU)
D) Dietary Orders — Dietary Orders will be sent to CIS in delimited form when available. Including but not limited to: supplemental feedings, tube feedings, NPO status.
Purpose: Dietary orders will be sent to provide treatment schedules in CIS.
a. Dietary orders will be sent to CIS for read-only use.
b. All order modifications will be made in VistA and contemporaneously displayed in CIS. Changes must be displayed in CIS both as a 'new' or 'critical' alert and in an updating field.
4.4.4 |  |  | D) Radiology Orders — All radiology orders will be sent to CIS. | Mandatory(ICU)
4.4.4.1 |  |  | 1. Purpose: radiology orders will be sent to facilitate nursing schedules in CIS. | Mandatory(ICU)
4.4.4.3 |  |  | a. Radiology orders will be sent to CIS for read-only use. | Mandatory(ICU)
4.4.4.4 |  |  | b. All order modifications will be made in VistA and contemporaneously displayed in CIS. Changes must be displayed in CIS both as a 'new' or 'critical' alert and in an updating field. | Mandatory(ICU)
4.4.4.5 |  |  | Radiology Results — All verified Radiology messages will be sent to CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.4.6 |  |  | 1. Purpose: To send Radiology Impression and Report results to the CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.4.8 |  |  | a1. Radiology Impression and Report results will be sent from VistA Radiology package to CIS for display. No limitations can be placed on maximum number of radiology data (specific tests or numbers of test results). | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.4.9 |  |  | a2. Radiology Impression and Report, including entire procedure name, exam name, and modifiers, must be displayed. All parent / child reports must be included with each update. | Mandatory(ICU)
4.4.4.10 |  |  | d. Query capabilities initiated by CIS will mimic CPRS (date range, specific results, etc.) | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.4.11 |  |  | e. Changes must be displayed in CIS both as a 'new' or 'critical' data alert and in an updating data field. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.5 |  |  | E) Vital Signs — Vital signs will be sent to ARK or CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.5.1 |  |  | 1. Purpose: A record of the complete patient vital sign information from VISTA (including O2 saturation and pain scores, as available) will be sent to the CIS.. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.5.3 |  |  | a. Vital sign data will be sent to the ARK or CIS for use by clinicians. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.5.4 |  |  | b. Updated data will be sent to the CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
|  |  | c. Height and Weight from VISTA will be sent to ARK and CIS. If no height or weight is available in VISTA each can be entered manually in CIS or ARK. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.6 |  |  | F) Problem List — Current Problem List entries will to be sent to CIS or ARK. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.6.1 |  |  | 1. Purpose: The Problem List contain information on current patient problems that may be valuable in assessing patients. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.6.3 |  |  | a. The problem list is sent initiated by the standard triggers and modifications to the list. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.6.4 |  |  | b. ICD-9/ICD-10 codes will be included. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.7 |  | Make this section mandatory across all systems | G) Notification of New Messages — A ‘Notification’ message will identify all new messages sent to CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.7.1 |  |  | 1. Purpose: The Notification message will be used by the ARk or CIS to alert the clinical user of new messages received by the ARK/CIS. The ARK/CIS will display the message type in efficient and easily comprehensible methods for clinical users (e.g., color coded icon indicating 'new labs', 'out of range labs', 'new orders'). | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.7.3 |  |  | a. When a new message is created, a second, ‘Notification’ message is sent also to the ARK/CIS identifying alerts, new lab values, etc. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.7.4 |  |  | b. The ARK/CIS will use the information in the Notification message to inform the clinical user of receipt of new data. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.4.7.5 |  | Make this section mandatory across all systems | c. The ARK/CIS will display a 'connection interrupted' message obvious to the clinical provider within 5 minutes of a communication interruption ( VistA, DSS, or network failure) between the ARK/CIS and VistA. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.5 |  | Time Synchronization — The times in the CIS systems must be synchronized with VistA.
4.5.1 |  |  | 1. Purpose: To have an accurate medical record the ARK/ CIS clocks must be identical to the VistA time. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.5.2 |  |  | a. The ARK/CIS must receive a timing signal from the interface and the synchronize the ARK/ CIS time with the VistA time. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
b. Time zones must be synchronized with the database when time changes are made.
4.6 |  | VistA Imaging — A .pdf image is sent from CIS to VistA Imaging will be associated with a TIU note.
4.6.1 |  |  | 1. Purpose: Since dense data and many data types cannot be stored in VistA, the CIS must create an image file for permanent, medical record storage in VistA. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.6.2.1 |  |  | a. Images of the ICU flowchart, PACU record, and anesthesia record will be created and sent to VistA Imaging by the ARK/ CIS at both pre-determined and user-defined intervals for permanent display and storage. Names and credentials of all clinicians contributing to the documentation within the image must be included therein. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.6.2.2 |  |  | b. The images will be attached to a TIU note title and administratively signed. Must be able to attach provider signature, when required. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.6.2.3 |  |  | c. Similar to paper flow sheets and records, the images display the initials of the individual who entered a particular value. The individual’s full name is displayed as a legend on the form. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
|  |  | d. All edit trail will be visiable in the application and included in the PDF . | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.6.2.4 |  |  | e. PDF and TIU is the required format. A national format should be used when available. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
|  |  | f. Any documentation in the application should reflect the name of the documentor in the application and the PDF. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.7 |  | TIU (Progress Note) — A progress note is created in the CIS and sent to VistA.
4.7.1 |  |  | 1. Purpose: A TIU note signed by the clinician is created in CIS and/or ARK and placed in CPRS Progress Notes. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.7.2.1 |  |  | a. Progress notes created within CIS and/or ARK will be exported to CPRS for permanent display and storage in VistA Progress Notes. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.7.2.2 |  |  | b. The note may be reviewed, edited and electronically signed in VistA. It may also be signed in the CIS or administratively signed in VistA. The capability for both approaches are required in the CIS. D229 | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.7.2.3 |  |  | c. The ARK/CIS must have the capacity for “all documented elements” and “only currently documented elements” in all flow sheets to be included in the progress note automatically and these documented elements in their entire meaning must be available and configurable by the site. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.7.2.4 |  |  | d. Electronic signatures applied in ARK/CIS along with the signed progress note will be transmitted to VistA. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
4.7.2.5 |  |  | e. Electronic signature information will be accessed/validated in VistA at the time of signing into ARK/CIS. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
Each configured note type must be transferable to a TIU note by manual means (click a ‘send’ icon) and/or automatically on a configurable, recurring time basis and/or when triggered by an event such as admission, discharge, etc.
4.8 |  | Anesthesia Record
4.8.1 |  |  | The ARK must display to the anesthesia provider all physiologic data received by the ARK within an acceptable time (less than 10 seconds) following the initiation of data reception. |  |  | Mandatory (ARK)
4.8.2 |  |  | Although the entire data stream must be captured in the ARK database, recording of patient physiologic data in the final anesthesia record must be performed at an interval no greater than every 1 minute. The initial set of vitals signs, however, must be time-stamped and documented in the final anesthesia record. |  |  | Mandatory (ARK)
The data stream from patient monitors must use serial connections and not network data flow.
4.8.3 |  |  | Upon completion of the anesthesia procedure, the ARK record must remain open to annotations and additions for a facility-specific, predetermined length of time, after which the record is closed and exported to VistA imaging. This time period may extend up to 48 hours. Once closed and exported for permanent storage, further annotations to the record must be prohibited. |  |  | Mandatory (ARK)
4.8.4 |  |  | The VistA electronic signature (or co signature) of the anesthesia provider must authenticate the completed anesthesia record and any VistA Progress Notes. The name and credentials of the authenticating anesthesia provider must appear in the completed electronic anesthesia record, whether displayed as a hard copy, an electronic Portable Document Format (PDF) file or an electronic display. |  |  | Mandatory (ARK)
4.8.5 |  |  | The completed electronic anesthesia record, whether displayed as a hard copy or an electronic Portable Document Format (PDF) file, must include the physiologic/pharmacologic record of intraoperative events with its attendant user annotations. It must also include the patient’s full name, social security number, age, pre-anesthesia American Society of Anesthesiologists (ASA) physical classification, and allergies, and the names of all anesthesia personnel involved in the procedure, the date of the procedure, the anesthesia and surgical start and stop times, and the principal type of anesthesia employed. All perioperative procedures and the administration of peri-operative antibiotics must be clearly identified. All anesthesia procedures must be clearly identified. |  |  | Mandatory (ARK)
4.8.6 |  |  | The completed electronic anesthesia record must contain documentation of the passing of patient care from the operating room anesthesia provider to the Post-Anesthesia Care Unit (PACU) provider. This documentation must include the patient’s blood pressure, heart rate, oxygen saturation, respiratory rate, and temperature upon arrival in the PACU, the patient’s method of oxygenation, and confirmation that all pertinent patient care information has been exchanged with the PACU provider. |  |  | Mandatory (ARK)
4.9 |  | Surgery Package Data and Schedule — The surgery schedule must provide scheduling information as well as patient demographics and other information (see Section II) to the CIS and/or ARK. |  |  |  | Mandatory (ARK)
4.9.1 |  |  | A) Surgery Schedule — All configured HL-7 message defined by Data bridge are sent to ARK/PACU CIS. The surgery scheduling case messages consist of case creation, case rescheduling, case modification and case cancellation. |  |  | Mandatory (ARK)
4.9.1.1 |  |  | 1. Purpose: Provide the patient lists in the surgery schedule to the ARK. |  |  | Mandatory (ARK)
4.9.1.3 |  |  | a. Additional data fields in the Surgery Package will be exchanged between the ARK/PACU CIS and the Surgery Package. These are included in the Registration information and other sources such as Labs. |  |  | Mandatory (ARK)
4.9.2 |  |  | B) Surgery Package Data — Data are exchanged between the ARK and surgery package. |  |  | Mandatory (ARK)
4.9.2.1 |  |  | 1. Purpose: There are multiple fields that are shared between the surgery package and the ARK. |  |  | Mandatory (ARK)
4.9.2.3 |  |  | a. The following fields have been identified as necessary areas that must allow for modification by the clinical user: |  |  | Mandatory (ARK)
4.9.2.4 |  |  | i. Time Patient In OR |  |  | Mandatory (ARK)
4.9.2.5 |  |  | ii. Time Patient Out OR |  |  | Mandatory (ARK)
4.9.2.6 |  |  | iii. Time Operation Began and Time Operation End |  |  | Mandatory (ARK)
4.9.2.7 |  |  | iv. Anesthesia Care Start and Anesthesia Care End |  |  | Mandatory (ARK)
v. Induction Complete
4.9.2.9 |  |  | vi. Principal Anesthetist |  |  | Mandatory (ARK)
4.9.2.10 |  |  | vii. Supervising Anesthesiologist |  |  | Mandatory (ARK)
4.9.2.11 |  |  | viii. Post Op BP, Temperature, Pulse and Respiratory Rate |  |  | Mandatory (ARK)
4.9.2.12 |  |  | ix. ASA Class |  |  | Mandatory (ARK)
4.9.2.13 |  |  | x. Other Surgery messages: the ORU messages are sent during the case updating the CIS with over 100 values entered into Vista during the case. Examples are ASA score, blood loss, body fluid measurements, etc. |  |  | Mandatory (ARK)
4.9.2.14 |  |  | xi. A means to reconcile data between the Surgery Package and ARK will be provided. |  |  | Mandatory (ARK)
4.9.2.15 |  |  | b. The ARK gives a visual representation of the area(s) in the VistA Surgical Package that are currently not updated for the surgery that was performed. |  |  | Mandatory (ARK)
Provider must have the ability to free text the procedure/surgical procedure into the procedure performed field.
All information relevant to the billing of anesthesia professional services must be captured by the CIS/ARK
| ARK Billing Requirements |  | VistA fields that contain information relevant to the billing of anesthesia professional services must be populated by the CIS/ARK via HL7 protocol pathways. This information includes, but is not limited to American Society of Anesthesiologists (ASA) physical classification, the names of all anesthesia personnel involved in the procedure, the date of the procedure, the anesthesia and surgical start and stop times, and the principle type of anesthesia employed.
The CIS/ARK report of anesthesia professional services must be permanently stored in a local database, in a fashion that is easily accessible for audit.
The CIS/ARK report of anesthesia professional services must be permanently stored in a local database, in a fashion that is easily accessible for audit.
5.0 | Additional CIS Clinical Assessment Capabilities |  | THE FOLLOWING CLINICAL ASSESSMENT APPLICATIONS REQUIRE A RESPONSE AS TO WHETHER YOUR SYSTEM HAS THE CAPACITY TO ACCOMPLISH THE DESCRIBED ACTION. Describe in detail and/or provide screen capture images demonstrating how your application accomplishes the desired action(s).
5.1 |  | Document body system assessments |  | Mandatory(ICU) | Mandatory(PACU) |  | Mandatory (Pre-Op)
5.1.1 |  |  | The capability to document body system assessments sufficient to provide the Unit with a detailed head-to-toe assessment that is configurable to each Unit, level of care, facility and groups of patients e.g. liver transplant or open heart. | Mandatory(ICU) | Mandatory(PACU) |  | Mandatory (Pre-Op)
5.1.2 |  |  | CIS and/or ARK must have the capacity to create a TIU for transfer note documentation between units/services must include required elements as defined by the VA. Documentation shall include but not be limited to the following: most recent vitals, prior 24 hour Intake and output, most recent assessment, MRSA swab, current IV fluids and medications, IV status, continuous IV fluids, IV data and time on lines. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.1.3 |  |  | The CIS and/or ARK must include severity scales. Severity scales must be calculated automatically (with mandatory manual data elements, as required), such as Apache II, III and IV, APS, Braden, Aldrete, Glascow coma scale, Morse fall scale, nutrition scoring, etc. as well as locally created scores. | Mandatory(ICU) | Mandatory(PACU)
5.2 |  | Manage calculations and calculated data |  | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
5.2.1 |  |  | The ability to accept calculated values from other devices (including data devices such as monitors etc.). | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
5.2.2 |  |  | The ability to perform manual and automatic clinical calculations. The calculations should include but not be limited to the following: IV weight based IV medications, physiological calculations | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
5.3 |  | Enter data and conduct fluid balance management |  | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
5.3.1 |  |  | The ability of clinicians on different shifts to enter data into a system that is capable of distinguishing fluids (blood, drips, IV fluids, etc.) and manage continuous fluid balance calculations inclusive of unit stay. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
5.3.2 |  |  | The ability to identify net negative I/O. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
5.3.3 |  |  | The ability to document all blood product administration as individual components. (for fluid management) | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK)
|  |  | Continuous infusions should be capable of being displayed using units provided by the site. If an infusion is compounded by the pharmacy, nursing, or anesthesia, the final compounded concentration in selectable units must be displayed (mcg/ml, mcg/kg/min, mcg/kg/hr). | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.4 |  | Intuitive user interface |  | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.4.1 |  |  | The user inter face should be easy and implicit: colors, icons, graphs, pick lists, freeform text, notes, forms, point and click, drop and drag, and other normalized Word-type capabilities are critical to reducing errors. The size of text, numbers, graphs must be easily readable in a high or low light environment. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.4.2 |  |  | The user interface must be 508 compliant, where applicable. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5 |  | Charting Capabilities |  | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.1 |  |  | Site configurable fields and views according to unit. Compounded medications should be displayed in selectable units with the compounded final concentration. Bolus and infusion lables should be the administered concentration or dose. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.2 |  |  | Ability to mandate selection of entries from pick-lists to minimize free-text | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.3 |  |  | Ability to support TIU templates, procedures including but not limited to: H&P, patient problems. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.4 |  |  | A method must be in place that will allow the user to annotate artifactual data, with an audit trail of corrected data maintained. If the CIS has data-averaging capabilities, an audit trail of automatic corrections must also be maintained. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.5 |  |  | Ability to print and or transfer to VistA imagining a PDF or TIU q-shift and other summary reports | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.6 |  |  | Create, setup and manage administrations of infusions and complex drips . Final concentration of compounded medications in user selectable units (time and or weight base units) | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.7 |  |  | Ability to auto calculate I’s and O’s (data displayed by tables, variable time periods, graphically, separated into colloid, crystalloid, blood // urine, NG, chest tube, etc.) | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.8 |  |  | Display tabular and graphical values (including from CPRS with vital signs, laboratory values, and other data)—all on the same timeline | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.9 |  |  | The ability to display I/O balance (preferably graphically) over a user selected time period. | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
5.5.10 |  |  | Ability to easily configure the flowchart to the individual patient, i.e., add or subtract rows as patient condition changes | Mandatory(ICU) | Mandatory(PACU) | Mandatory (ARK) | Mandatory (Pre-Op)
THIS IS THE MAIN EHRM-SPECIFIC REQUIREMENT; ALL OTHERS CIS/ARK REQUIREMENTS ABOVE ARE ALREADY SATISFIED FOR VISTA
6.0 | Additional CIS/ARK Interfaces (Cerner / EHRM) |  | The folliwng additional interfaces are required to maintain seamless continuity of care and services for all veterans, maintaining the full functaionality of all existing CIS/ARK systems, and interfacing these CIS/ARK systems to the Cerner (EHRM) system to the same standards they are interfaced to VistA.
6.1 |  | Vendor-side interface (see section 4.1) | VHA CIS/ARK vendors shall interface to EHRM using the same vendor-side interfaces as specified in section 4.1. (Vendor-side interface). This will allow VHA anesthesiology staff may continue to use existing CIS/ARK for all anesthesia functionality, but use the EHRM (Powerchart) for all other charting functions instead of CPRS. The two main CIS/ARK vendors shall provide, test, validate, and operationalize this Vendor-side interface to EHRM: (a) Picis (2) Draeger (3) others as required